## Data
UCT AST9AH Lead Acid Battery Low Temperature Testing
March 2024
Document Information
Analyst | Masamkele Mdondolo
Company | University of Cape Town
Attention | Lawrence Stanton
Tell | 071 493-6106
Email | STNLAW003@myuct.ac.za
Postal | 4th Floor Menzies Building, Upper Campus, University of Cape Town, Cape Town, 7701
Methodology | UCT SISPS Low Temperature Tests Methodology 2024.md
Preparatory Work
Due Date | Responsible | Task
2024-03-01 | Masa | Exploratory teardown of defective battery. Photograph internal construction.
2024-03-01 | Masa | Verify operation of ETC to -40°C (empty).
2024-03-01 | Lawrence | Construct EIS clamp rig.
2024-03-01 | Lawrence | Label all samples and batteries.
2024-03-02 | Lawrence | Load batteries and sample bottles for transit.
2024-03-04 | Masa | Assemble lab apparatus: Cutting tool, filter funnel, acetone, PPE, pipettes.
2024-03-04 | Masa | Discharge, drill into, and place thermometer into sacrificial battery.
2024-03-04 | Masa | Prepare battery cycling program.
2024-03-04 | Masa | Verify EIS test program on dummy battery.
1. Cycling
Done
Scheduled Start: | 2024-03-04 12:00 | No
Scheduled End: | 2024-03-12 12:00 | No
Make | Battery | Initial Voltage | Final Voltage | Change | Initial Weight | Final Weight | Difference | Channel |  | Log File Name | Notes / Interruptions
Asterion HR12-9 | UCT_AST9AH_A01 | 0 | 0 | 0 | 2.7802 |  | -100% | NA |  | NA | DO NOT CYCLE
Asterion HR12-9 | UCT_AST9AH_A02 |  |  | 0 | 2.7762 |  | -100% | NA |  | NA | DO NOT CYCLE
Asterion HR12-9 | UCT_AST9AH_A03 |  |  | 0 | 2.7617 |  | -100% | NA |  | NA | DO NOT CYCLE
Asterion HR12-9 | UCT_AST9AH_A04 |  |  | 0 | 2.7688 |  | -100% | NA |  | NA | DO NOT CYCLE
Asterion HR12-9 | UCT_AST9AH_A05 |  |  | 0 | 2.836 |  | -100% | NA |  | NA | DO NOT CYCLE
Asterion HR12-9 | UCT_AST9AH_A06 |  |  | 0 | 2.7717 |  | -100% | NA |  | NA | DO NOT CYCLE
Asterion HR12-9 | UCT_AST9AH_A07 |  |  | 0 | 2.7771 |  | -100% | NA |  | NA | DO NOT CYCLE
Asterion HR12-9 | UCT_AST9AH_A08 |  |  | 0 | 2.7641 |  | -100% | NA |  | NA | DO NOT CYCLE
Asterion HR12-9 | UCT_AST9AH_A09 |  |  | 0 | 2.7535 |  | -100% | NA |  | NA | DO NOT CYCLE
Asterion HR12-9 | UCT_AST9AH_A10 |  |  | 0 | 2.781 |  | -100% | NA |  | NA | DO NOT CYCLE
Asterion HR12-9 | UCT_AST9AH_B01 |  |  | 0 | 2.7748 |  | -100% | 7
Asterion HR12-9 | UCT_AST9AH_B02 |  |  | 0 | 2.7734 |  | -100% | 8
Asterion HR12-9 | UCT_AST9AH_B03 |  |  | 0 | 2.7838 |  | -100% | 12
Asterion HR12-9 | UCT_AST9AH_B04 |  |  | 0 | 2.7556 |  | -100% | 14
Asterion HR12-9 | UCT_AST9AH_B05 |  |  | 0 | 2.7665 |  | -100% | 15
Asterion HR12-9 | UCT_AST9AH_B06 |  |  | 0 | 2.768 |  | -100%
Asterion HR12-9 | UCT_AST9AH_B07 |  |  | 0 | 2.7915 |  | -100%
Asterion HR12-9 | UCT_AST9AH_B08 |  |  | 0 | 2.7893 |  | -100%
Asterion HR12-9 | UCT_AST9AH_B09 |  |  | 0 | 2.7851 |  | -100%
Asterion HR12-9 | UCT_AST9AH_B10 |  |  | 0 | 2.782 |  | -100%
Refer to methodology for cycling program and other battery limits.
2. Varied Discharge
Scheduled Start: |  |  | Done
| Batch A | 2024-03-04 11:30 | No
| Batch B | 2024-03-12 12:00 | No
Scheduled End:
| Batch A | 2024-03-04 16:50 | No
| Batch B | 2024-03-12 16:50 | No
Battery | Discharge Rate | Discharge Time | Discharge Time | Expected SoC | Channel | Log File Name | Notes / Interruptions
UCT_AST9AH_A01 | NA | 0 | 0 | 100% | 15 | UCT-AST9AH-A01.csv | Used for initial EIS tests at factory SoC. Then ran through same Bulk-Absorption-Float charge as other batteries, except zero discharge.
UCT_AST9AH_A02 | 0.9 | 0.027778 | 40 | 93.33% | 14 | UCT-AST9AH-A02.csv
UCT_AST9AH_A03 | 0.9 | 0.055556 | 80 | 86.67% | 13 | UCT-AST9AH-A03.csv
UCT_AST9AH_A04 | 0.9 | 0.083333 | 120 | 80% | 15 | UCT-AST9AH-A04.csv
UCT_AST9AH_A05 | 0.9 | 0.11111 | 160 | 73.33% | 15 | UCT-AST9AH-A05.csv | Interrupted in Absorption to increase cut-off from 0.2A to 0.5A (taking too long). Rapidly stopped test and resumed in new routine, skipping to Absorption step. Second run is "-X".
UCT_AST9AH_A06 | 0.9 | 0.13889 | 200 | 66.67% | 14 | UCT-AST9AH-A06.csv | Interrupted in Absorption to increase cut-off from 0.2A to 0.5A (taking too long). Rapidly stopped test and resumed in new routine, skipping to Absorption step. Second run is "-X".
UCT_AST9AH_A07 | 0.9 | 0.16667 | 240 | 60% | 13 | UCT-AST9AH-A07.csv | Interrupted in Absorption to increase cut-off from 0.2A to 0.5A (taking too long). Rapidly stopped test and resumed in new routine, skipping to Absorption step. Second run is "-X".
UCT_AST9AH_A08 | 0.9 | 0.19444 | 280 | 53.33% | 12 | UCT-AST9AH-A08.csv | Interrupted in Absorption to increase cut-off from 0.2A to 0.5A (taking too long). Rapidly stopped test and resumed in new routine, skipping to Absorption step. Second run is "-X".
UCT_AST9AH_A09 | 0.9 | 0.22222 | 320 | 46.67% | 7 | UCT-AST9AH-A09.csv | Interrupted in Absorption to increase cut-off from 0.2A to 0.5A (taking too long). Rapidly stopped test and resumed in new routine, skipping to Absorption step. Second run is "-X".
UCT_AST9AH_A10 | 0.9 | 0.25 | 360 | 40% | 8 | UCT-AST9AH-A010.csv | Interrupted in Absorption to increase cut-off from 0.2A to 0.5A (taking too long). Rapidly stopped test and resumed in new routine, skipping to Absorption step. Second run is "-X". Initially programmed for 11V cut-off assuming 9Ah capacity. Note below details how capacity appears reduced. Expected 11.0V threshold to hit, halting test. Opted to remote into terminal later and manually resume test, manually monitoring voltage until end of 6h discharge.
UCT_AST9AH_B01 | 0.9 | 0 | 0 | 100% | 7
UCT_AST9AH_B02 | 0.9 | 0.027778 | 40 | 93.33% | 8
UCT_AST9AH_B03 | 0.9 | 0.055556 | 80 | 86.67% | 12
UCT_AST9AH_B04 | 0.9 | 0.083333 | 120 | 80% | 14
UCT_AST9AH_B05 | 0.9 | 0.11111 | 160 | 73.33% | 15
UCT_AST9AH_B06 | 0.9 | 0.13889 | 200 | 66.67%
UCT_AST9AH_B07 | 0.9 | 0.16667 | 240 | 60%
UCT_AST9AH_B08 | 0.9 | 0.19444 | 280 | 53.33%
UCT_AST9AH_B09 | 0.9 | 0.22222 | 320 | 46.67%
UCT_AST9AH_B10 | 0.9 | 0.25 | 360 | 40%
Deeply discharged batteries can be left to finish after working hours as those are only added to ETC at higher temperature the next day.
Observation: All batteries were found to be fairly discharged, assumed due to self discharge following storage period.
Adjustment: Increased Absorption cut-off current from 0.2A to 0.5A.
Observation: At 1h20 into discharge, batteries presented a discharge curve suggesting <9Ah capacity. Datasheet inspection showed slightly derated nominal capacity, 8.6Ah at 0.86A discharge. Calculated 5.6Ah capacity based off discharge curve and 10.5V 100% DoD before leaving for the day.
3. Cooling & EIS
Min Time Spacing: | 0:30
Max Temp Deviation: | 2
Schedule
| Time per EIS | 0:10
| Temperature Increment | 3:00
Battery | Target Temperature | Due | Temperature Eth Glycol | Reading Time | Battery Temperature Logic | Battery Min Temperature | Battery Max Temperature | EIS Start Time | EIS Log File | Measurement Notes / Interruptions | Temperature Notes
UCT_AST9AH_A01 | 25 | 2024-03-04 11:30 |  |  |  | 25 | 28 |  | UCT_AST_9AH_A01_+RT1 | Performed several test EIS measurements at SoC on arrival. Then recharged before final EIS made. Measurement of A03-A01 performed after A10-A04 due to sequencing on battery tester. | Temperature range assumed between ambiant temperatures of the 2 rooms involved.
UCT_AST9AH_A02 | 25 | 2024-03-04 12:10 |  |  |  | 25 | 28 |  | UCT_AST_9AH_A02_+RT1 | Left to rest for 30min-1h before measurement completed.
UCT_AST9AH_A03 | 25 | 2024-03-04 12:50 |  |  |  | 25 | 28 |  | UCT_AST_9AH_A03_+RT1 | Measurement started soon after discharge completed.
UCT_AST9AH_A04 | 25 | 2024-03-04 13:30 |  |  |  | 25 | 28 |  | UCT_AST_9AH_A04_+RT1 | Left to rest for several hours after finishing discharge overnight
UCT_AST9AH_A05 | 25 | 2024-03-04 14:10 |  |  |  | 25 | 28 |  | UCT_AST_9AH_A05_+RT1 | Left to rest for several hours after finishing discharge overnight.
UCT_AST9AH_A06 | 25 | 2024-03-04 14:50 |  |  |  | 25 | 28 |  | UCT_AST_9AH_A06_+RT1 | Left to rest for several hours after finishing discharge overnight.
UCT_AST9AH_A07 | 25 | 2024-03-04 15:30 |  |  |  | 25 | 28 |  | UCT_AST_9AH_A07_+RT1 | Left to rest for several hours after finishing discharge overnight.
UCT_AST9AH_A08 | 25 | 2024-03-04 16:10 |  |  |  | 25 | 28 |  | UCT_AST_9AH_A08_+RT1 | Left to rest for several hours after finishing discharge overnight.
UCT_AST9AH_A09 | 25 | 2024-03-05 09:00 |  |  |  | 25 | 28 |  | UCT_AST_9AH_A09_+RT1 | Left to rest for several hours after finishing discharge overnight.
UCT_AST9AH_A10 | 25 | 2024-03-05 09:10 |  |  |  | 25 | 28 |  | UCT_AST_9AH_A10_+RT1 | Left to rest for several hours after finishing discharge overnight.
UCT_AST9AH_B01 | 25 | [8 empty] | Used on Wed 6 March to debug EIS issues while still at factory SoC.
UCT_AST9AH_B02 | 25
UCT_AST9AH_B03 | 25
UCT_AST9AH_B04 | 25
UCT_AST9AH_B05 | 25
UCT_AST9AH_B06 | 25
UCT_AST9AH_B07 | 25
UCT_AST9AH_B08 | 25
UCT_AST9AH_B09 | 25
UCT_AST9AH_B10 | 25
UCT_AST9AH_A01 | -40 | 2024-03-05 09:30 | -38.8 | 9:35 | Between A01 and A06. | -38.8 | -32.4 | 9:33 | UCT_AST_9AH_A01_-40 | Nominal. Photo taken of ice deposition near terminals, but otherwise well insulated. Transferred A01-A05 from ETC to freezer as this measurement was started. Eth Glycol left in ETC with A06-A08. | Temperature measurement is of Eth Glycol made shortly after transferring A01-A05 from ETC to Freezer, leaving A06-A08 in ETC and returning Eth Glycol to ETC. This behaviour was later altered after observing ETC behaviour. Secondary freezer generally reporting -30°C while performing EIS measurements for A01-A05.
UCT_AST9AH_A02 | -40 | 2024-03-05 09:40 |  |  | Between A01 and A06. | -38.8 | -32.4 | 9:43 | UCT_AST_9AH_A02_-40 | Nominal.
UCT_AST9AH_A03 | -40 | 2024-03-05 09:50 |  |  | Between A01 and A06. | -38.8 | -32.4 | 9:53 | UCT_AST_9AH_A03_-40 | Nominal.
UCT_AST9AH_A04 | -40 | 2024-03-05 10:00 |  |  | Between A01 and A06. | -38.8 | -32.4 | 10:53 | UCT_AST_9AH_A04_-40 | Nominal.
UCT_AST9AH_A05 | -40 | 2024-03-05 10:10 |  |  | Between A01 and A06. | -38.8 | -32.4 | 10:13 | UCT_AST_9AH_A05_-40 | Nominal.
UCT_AST9AH_A06 | -40 | 2024-03-05 10:20 | -32.4 | 10:20 | Below A08. | -38.8 | -27.8 | 10:23 | UCT_AST_9AH_A06_-40 | Nominal. Transferred A01-A05 from freezer back to ETS, and A06-A08 and Eth Glycol from ETC to freezer. | ETC drifted to max -25°C before beginning recovery to -40°C setpoint for A06-A08. Temperature measured when transferring A06-A08 from ETC to freezer following completion of A01-A05 EIS. Eth Glycol then placed in freezer for remainder of this temperature step.
UCT_AST9AH_A07 | -40 | 2024-03-05 10:30 |  |  | Below A08. | -38.8 | -27.8 | 10:33 | UCT_AST_9AH_A07_-40 | Transferred A06-A09 and Eth Glycol from freezer to ETC following this measurment.
UCT_AST9AH_A08 | -40 | 2024-03-05 10:40 | -27.8 | 10:40 | NA |  |  |  | NOT MEASURED | Measurement cancelled. Forgot before initiating transfer of all samples. Eth Glycol measurement however indicates significant drift form -40°C, so would be quite invalid at this stage regardless.
UCT_AST9AH_A09 | NA | [7 empty] | NOT MEASURED
UCT_AST9AH_A10 | NA | [7 empty] | NOT MEASURED
UCT_AST9AH_B01 | -40 |  | -39.3 | 12:00 | Between B01 and B06 | -39.3 | -30.9 | 12:00 | UCT_AST_B01_-40 | Nominal | B01-B08 placed at -40°C in ETC from Friday 10 May at ±15:00 and left over weekend until begin of work on 13 May 12:00. Then transferred into chest freezer. Initial temperature measured with accompaning EthGly solution.
UCT_AST9AH_B02 | -40 |  |  |  | Between B01 and B06 | -39.3 | -30.9 | 12:12 | UCT_AST_B02_-40
UCT_AST9AH_B03 | -40 |  |  |  | Between B01 and B06 | -39.3 | -30.9 | 12:23 | UCT_AST_B03_-40
UCT_AST9AH_B04 | -40 |  |  |  | Between B01 and B06 | -39.3 | -30.9 | 12:33 | UCT_AST_B04_-40 |  | Added fan to cool room.
UCT_AST9AH_B05 | -40 |  |  |  | Between B01 and B08 | -39.3 | -28.6 | 13:57 | UCT_AST_B05_-40 | Anominly. Presumed cable fell down from bench while measurement occurred. Measurement repeated later after B08. Some curves cfaield with erranous data points. Repeated measurements (stopping some early) but no single spectra was clean. Accepted manual processing of final spectra discarding 5Hz.
UCT_AST9AH_B06 | -40 |  | -30.9 | 12:57 | Between B01 and B08 | -39.3 | -28.6 | 12:57 | UCT_AST_B06_-40
UCT_AST9AH_B07 | -40 |  |  |  | Between B01 and B08 | -39.3 | -28.6 | 13:20 | UCT_AST_B07_-40 | Initial measurement failed - program hanged withou submitting all data. Repeated measurement with success.
UCT_AST9AH_B08 | -40 |  | -28.6 | 13:30 | Between B01 and B08 | -39.3 | -28.6 | 13:30 | UCT_AST_B08_-40
UCT_AST9AH_B09 | -40 | [8 empty] | NOT MEASURED | B09-B10 placed at -30°C in chest freezer from 10 May 15:00 until 13 May 12:00 before swap into ETC. Remained in unit until B01-B08 returned to ETC for overnight -30°C setpoint.
UCT_AST9AH_B10 | -40 | [8 empty] | NOT MEASURED
UCT_AST9AH_A01 | -30 | 2024-03-05 15:30 | -34.3 | 13:31 | Between A01 and A06 | -34.3 | -29 | 15:31 | UCT_AST_9AH_A01_-30 | Freezer doors closed accedentially and held under vacuum. 30s delay in transferring batteries from ETC to freezer. A01 however proceeded nominally. | Following transfer of A01-A05 from freezer to ETC after completing -40°C tests, ETC reported drift up to -21°C before resuming cooling. Adjusted setpoint to -35°C temporarily expecting disturbance when A06-A08 will be returned to batch. Freezer set to -38°C to absorb expected thermal shock once EIS measurements begin. ETC achieved -35.0°C at 11:30, and therefore would be at this temoerature for 4h before EIS measurements begin.
UCT_AST9AH_A02 | -30 | 2024-03-05 15:40 |  |  | Between A01 and A06 | -34.3 | -29 | 15:40 | UCT_AST_9AH_A02_-30 | Nominal
UCT_AST9AH_A03 | -30 | 2024-03-05 15:50 |  |  | Between A01 and A06 | -34.3 | -29 | 15:59 | UCT_AST_9AH_A03_-30 | Positive lead disconnected without being noticed. EIS errored with same issue as encountered in morning. Required power chycle and repeated measurement attempts before resuming test. Total delay 9min before final EIS measurement started. | Any significant rise in temperature discounted due to insulation performance.
UCT_AST9AH_A04 | -30 | 2024-03-05 16:00 |  |  | Between A01 and A06 | -34.3 | -29 | 16:09 | UCT_AST_9AH_A04_-30 | Nominal
UCT_AST9AH_A05 | -30 | 2024-03-05 16:10 |  |  | Between A01 and A06 | -34.3 | -29 | 16:18 | UCT_AST_9AH_A05_-30 | Nominal
UCT_AST9AH_A06 | -30 | 2024-03-05 16:20 | -29 | 16:28 | Between A01 and A06 | -34.3 | -29 | 16:28 | UCT_AST_9AH_A06_-30 | Nominal | Following transfer of A06-A08 to ETC (whole batch reunited) ETC reported drift up to -15.2°C. Will maintain -35°C setpoint to compensate for 2h until 13:00, then hold at -32°C for 2h30 until 15:30 when EIS measurements will begin. Eth Glycol measurement performed as A06 removed for EIS.
UCT_AST9AH_A07 | -30 | 2024-03-05 16:30 |  |  | Between A01 and A09 | -34.3 | -27 | 16:37 | UCT_AST_9AH_A07_-30 | Nominal
UCT_AST9AH_A08 | -30 | 2024-03-05 16:40 |  |  | Between A01 and A09 | -34.3 | -27 | 16:47 | UCT_AST_9AH_A08_-30 | Nominal
UCT_AST9AH_A09 | -30 | 2024-03-05 16:50 | -27 | 16:57 | Between A01 and A09 | -34.3 | -27 | 17:04 | UCT_AST_9AH_A09_-30 | Ran POC EIS after leaving in secondary freezer overnight at -30°C. PoC EIS failed with error in programming VacMin. Decreased parameter value incrementally to 9V before deleting limit from test program (was optional), without success. Ran test on B1 at room temperature with identical error, suggesting issue with instrument. Repeated measurement on B1 but now succeeded (maybe issue in cooling the EIS leads?) Then tested again on A9 (after being left in Styrofoam holder for 30min) - DC voltage observed to increase from 10V5 to 12V1 and EIS measurement succeeded. Therefore inconclusive as to source of the original error. Actual measurement: EIS produced errenous values and stopped mid-test. Cleaned uce off termals and restarted test. | Placed in freezer with -38°C setpoint while completing -40°C step, assumin no significant cooling would occur during disturbances.
UCT_AST9AH_A10 | [8 empty] | NOT MEASURED
UCT_AST9AH_B01 | -30 | 2024-05-14 08:10 | -29.6 | 8:02 | Between B01 and B06 | -29.6 | -29.4 | 8:10 | UCT_AST_B01_-30 | Nominal | Batteries left in ETC overnight, no anomilies. Eth Glyol measurement taken upon opening, folloiwing which all samples and Eth Glycol transferred to chest freezer.
UCT_AST9AH_B02 | -30 | 2024-05-14 08:20 |  |  | Between B01 and B06 | -29.6 | -29.4 | 8:21 | UCT_AST_B02_-30 | Nominal
UCT_AST9AH_B03 | -30 | 2024-05-14 08:30 |  |  | Between B01 and B06 | -29.6 | -29.4 | 8:32 | UCT_AST_B03_-30 | Nominal
UCT_AST9AH_B04 | -30 | 2024-05-14 08:40 |  |  | Between B01 and B06 | -29.6 | -29.4 | 8:42 | UCT_AST_B04_-30 | Nominal
UCT_AST9AH_B05 | -30 | 2024-05-14 08:50 |  |  | Between B01 and B06 | -29.6 | -29.4 | 8:53 | UCT_AST_B05_-30 | Nominal
UCT_AST9AH_B06 | -30 | 2024-05-14 09:00 | -29.4 | 9:04 | Between B01 and B09 | -29.6 | -28.5 | 9:04 | UCT_AST_B06_-30 | Nominal
UCT_AST9AH_B07 | -30 | 2024-05-14 09:10 |  |  | Between B01 and B09 | -29.6 | -28.5 | 9:15 | UCT_AST_B07_-30 | Nominal
UCT_AST9AH_B08 | -30 | 2024-05-14 09:20 |  |  | Between B01 and B09 | -29.6 | -28.5 | 9:25 | UCT_AST_B08_-30 | Nominal
UCT_AST9AH_B09 | -30 | 2024-05-14 09:30 | -28.5 | 10:10 | [5 empty] | Measurement failed due to minimum voltage threshold. A cell appears to be collapsing. Attempted to adjust program without success. Probed terminal with multimeter but voltage holds near 11V5, without noticable drop when test program ran. After 20min the measurement was abandoned. Ran B02X afterwards (UCT_AST_B02X_-30) with success, elimating any possible issue with EIS instrument. | All batteries moved back to ETC for -20°C setpoint at 10:10 following abandonment of B09.
UCT_AST9AH_B10 | -30 | 2024-05-14 09:40 | [7 empty] | NOT MEASURED
UCT_AST9AH_A01 | -20 | 2024-03-07 07:45 | -23.2 | 7:52 | Between A06 and A01 | -24.6 | -23.2 | 7:56 | UCT_AST_9AH_A01_-20 | ETC found in morning to have maintained -25°C. Tests therefore began immediately. First EIS produced unstable values in HF range. Stopped and started new test. Error was then resolved. Second test nominal.
UCT_AST9AH_A02 | -20 | 2024-03-07 07:55 |  |  | Between A06 and A01 | -24.6 | -23.2 | 8:06 | UCT_AST_9AH_A02_-20 | Increased EIS program limit to 5kHz due to loss of real crossing in A09 @ -30°C. Nominal
UCT_AST9AH_A03 | -20 | 2024-03-07 08:05 |  |  | Between A06 and A01 | -24.6 | -23.2 | 8:18 | UCT_AST_9AH_A03_-20 | Nominal
UCT_AST9AH_A04 | -20 | 2024-03-07 08:15 |  |  | Between A06 and A01 | -24.6 | -23.2 | 8:28 | UCT_AST_9AH_A04_-20 | Nominal
UCT_AST9AH_A05 | -20 | 2024-03-07 08:25 |  |  | Between A06 and A01 | -24.6 | -23.2 | 8:38 | UCT_AST_9AH_A05_-20 | Nominal
UCT_AST9AH_A06 | -20 | 2024-03-07 08:35 | -24.6 | 8:48 | Between A06 and A01 | -24.6 | -23.2 | 8:48 | UCT_AST_9AH_A06_-20 | Placed A06 under test as Ethyl Glycol simultaniously removed for measuremnt. Nominal | Feezer came from -35°C shortly before being set to -25°C for this step. Decrease in temperature assumed due to this latent heat. Batteries therefore assumed to have possibly cooled within this range.
UCT_AST9AH_A07 | -20 | 2024-03-07 08:45 |  |  | Between A06 and A10 | -24.6 | -19.6 | 8:59 | UCT_AST_9AH_A07_-20 | Significant increase in impedance noticed in LF range. Otherwise Nominal.
UCT_AST9AH_A08 | -20 | 2024-03-07 08:55 |  |  | Between A06 and A10 | -24.6 | -19.6 | 9:09 | UCT_AST_9AH_A08_-20 | Nominal
UCT_AST9AH_A09 | -20 | 2024-03-07 09:05 |  |  | Between A06 and A10 | -24.6 | -19.6 | 9:19 | UCT_AST_9AH_A09_-20 | Noticable increase in impedance, suggesting damage from previous freezing observed at -30°C, or partial freezing still. Real crossing also lost, increased program limit from 5kHz to 8kHz to compensate. Measurement nominal
UCT_AST9AH_A10 | -20 | 2024-03-07 09:15 | -19.6 | 10:12 | Between A06 and A10 * | -24.6 | -19.6 | 9:52 | UCT_AST_9AH_A10_-20 RUN 1 + UCT_AST_9AH_A10_-20 RUN 2 | First measurement failed, stopped and started new test (with new program at 8kHz). Second run also failed - adjusted leads, started fan (turned off when measuring temperature at A06) and ran EIS test third time, faulure remained. Cleared top of terminals from ice, test failed but due to undervoltage limit (9VDC) at 5kHz, suggesting cell collapse, although other measurements return to ±11V. Battery appears destroyed, unable to hold voltage under conduction. Dropped voltage limit to 7VDC to investigate. Test still produced errenous results. Allowed to run through to investigate LF region. Test filled with errors and completely different spectrum but otherwise completed. Repeated second measurement to verify trend. Measuremnts display interesting features: grouping of spectrs and completely different profile in general. | Left in chest freezer over Wed 6 March until being rejoined with batch at end of Wed 6 March to cool at -25°C overnight. During measurement was out in styrofoam for 21min before final tests began.
UCT_AST9AH_B01 | -20 |  | -19.6 | 15:10 | Between B01 and B07 | -19.6 | -19.6 |  |  | Reduced program limits to 5Hz-3kHz. Nominal | All batteries moved from ETC to chest freezer at point of Eth Glycol measurement and start of B01. Weather mildly humid - noticable ice formation on all battery terminals while under EIS.
UCT_AST9AH_B02 | -20 |  |  |  | Between B01 and B07 | -19.6 | -19.6 |  |  | Nominal
UCT_AST9AH_B03 | -20 |  |  |  | Between B01 and B07 | -19.6 | -19.6 |  |  | Nominal
UCT_AST9AH_B04 | -20 |  |  |  | Between B01 and B07 | -19.6 | -19.6 |  |  | Nominal
UCT_AST9AH_B05 | -20 |  |  |  | Between B01 and B07 | -19.6 | -19.6 |  |  | Measurement finished but 4x errenous data pints. Can be corrected but re-ran test, citing issue with the instrument and otherwise valid curve. Re-run produced similar error at ±100Hz but otherwise improved.
UCT_AST9AH_B06 | -20 |  |  |  | Between B01 and B07 | -19.6 | -19.6 |  |  | Nominal
UCT_AST9AH_B07 | -20 |  | -19.6 | 16:15 | Between B01 and B07 | -19.6 | -19.6 |  |  | Nominal
UCT_AST9AH_B08 | -20 |  |  |  | Between B01 and B10 | -19.6 | -19 |  |  | Nominal
UCT_AST9AH_B09 | -20 |  |  |  | Between B01 and B10 | -19.6 | -19 |  |  | Nominal
UCT_AST9AH_B10 | -20 |  | -19 | 16:50 | Between B01 and B10 | -19.6 | -19 |  |  | Noninal - continous curve unlike A10 | Batteries transferred back to ETC following B10 EIS.
UCT_AST9AH_A01 | -10 | 2024-03-07 13:00 | -9.9 | 15:31 | Between A01 and A05 | -9.9 | -6.4 | 15:31 | UCT_AST_9AH_A01_-10 | 8kHz reading records very low Real Z. Deemed outlier and should be filtered out. Otherwise nominal. | Given previous setpoint, placed this measurement's saturation to -18°C for axis spacing. Realised however, due to delays in A10 at -20°C, that insufficient time remained. Therefore changed setpoint to -10°C and increased saturation time. Loadshedding struck at 12:00, chamber drifted to -4.3°C before resuming cooling rapidly back to setpoint. Freezer not available while running another test. Forced to use ETC for this temperature step. Did not attempt to operate between batteries (airflow calibration heats chamber).
UCT_AST9AH_A02 | -10 | 2024-03-07 13:10 |  |  | Between A01 and A05 | -9.9 | -6.4 | 15:42 | UCT_AST_9AH_A02_-10 | Similar to previous (8kHz measurement).
UCT_AST9AH_A03 | -10 | 2024-03-07 13:20 |  |  | Between A01 and A05 | -9.9 | -6.4 | 15:52 | UCT_AST_9AH_A03_-10 | Similar to previous (8kHz measurement). | ETC drift at -5°C.
UCT_AST9AH_A04 | -10 | 2024-03-07 13:30 |  |  | Between A01 and A05 | -9.9 | -6.4 | 16:03 | UCT_AST_9AH_A04_-10 | Similar to previous (8kHz measurement).
UCT_AST9AH_A05 | -10 | 2024-03-07 13:40 | -6.4 | 16.13 | Between A01 and A05 | -9.9 | -6.4 | 16:13 | UCT_AST_9AH_A05_-10 | Similar to previous (8kHz measurement). | Eth Glycol measurement taken as A05 removed for testing.
UCT_AST9AH_A06 | -10 | 2024-03-07 13:50 |  |  | Between A01 and A10 | -9.9 | -1 | 16:23 | UCT_AST_9AH_A06_-10 | Similar to previous (8kHz measurement). | ETC at -1.0°C.
UCT_AST9AH_A07 | -10 | 2024-03-07 14:00 |  |  | Between A01 and A10 | -9.9 | -1 | 16:34 | UCT_AST_9AH_A07_-10 | Similar to previous (8kHz measurement).
UCT_AST9AH_A08 | -10 | 2024-03-07 14:10 |  |  | Between A01 and A10 | -9.9 | -1 | 16:44 | UCT_AST_9AH_A08_-10 | Similar to previous (8kHz measurement).
UCT_AST9AH_A09 | -10 | 2024-03-07 14:20 |  |  | Between A01 and A10 | -9.9 | -1 | 16:54 | UCT_AST_9AH_A09_-10 | Similar to previous (8kHz measurement).
UCT_AST9AH_A10 | -10 | 2024-03-07 14:30 | -1 |  | Between A01 and A10 | -9.9 | -1 | 16:06 | UCT_AST_9AH_A10_-10 |  | Battery felt noticably colder than A09, possible still frozen.
UCT_AST9AH_B01 | -10 | 2024-05-15 08:00 | -9.8 |  | Between B01 and B06 | -9.8 | -8.4 | 8:10 | UCT_AST_B01_-10 | Nominal | Batteries moved from ETC to chest freezer simultanious with start of B01 and Eth Glycol reading. Freezer was set to -15°C, for unknown reasons. Adjusted setpoint to -8°C, shocked freezer by opening door, and observed stabilisation near -10°C after batteries transferred. Intervention regarded as successful in rapidly changing temperature to -10°C. Setpoint later resumed to -10°C.
UCT_AST9AH_B02 | -10 | 2024-05-15 08:10 |  |  | Between B01 and B06 | -9.8 | -8.4 | 8:20 | UCT_AST_B02_-10 | Nominal
UCT_AST9AH_B03 | -10 | 2024-05-15 08:20 |  |  | Between B01 and B06 | -9.8 | -8.4 | 8:31 | UCT_AST_B03_-10 | Nominal
UCT_AST9AH_B04 | -10 | 2024-05-15 08:30 |  |  | Between B01 and B06 | -9.8 | -8.4 | 8:41 | UCT_AST_B04_-10 | Nominal
UCT_AST9AH_B05 | -10 | 2024-05-15 08:40 |  |  | Between B01 and B06 | -9.8 | -8.4 | 8:51 | UCT_AST_B05_-10 | Nominal
UCT_AST9AH_B06 | -10 | 2024-05-15 08:50 | -8.4 |  | Between B01 and B06 | -9.8 | -8.4 | 9:01 | UCT_AST_B06_-10 | Nominal | B01-B06 transferred to ETC (set to 0°C before being unpowered after transfer) due to limited settling time available before 0°C tests. B07-B09 will be returned to chest freezer at -10°C howevr until B10 is complete.
UCT_AST9AH_B07 | -10 | 2024-05-15 09:00 |  |  | Between B01 and B06 | -9.8 | -8.4 | 9:22 | UCT_AST_B07_-10 | Double errenous readings in LF range. Elected to repeat EIS measurement. Second successful except for single error data point near 50Hz. Observations suggestmains power noise issue.
UCT_AST9AH_B08 | -10 | 2024-05-15 09:10 |  |  | Between B01 and B06 | -9.8 | -8.4 | 9:34 | UCT_AST_B08_-10 | Nominal
UCT_AST9AH_B09 | -10 | 2024-05-15 09:20 |  |  | Between B01 and B06 | -9.8 | -8.4 | 9:44 | UCT_AST_B09_-10 | Nominal | Chest freeze emptied with B06-B09 transferred. EIS remains unpowered pending B10. Chest freezer set to -5°C (maximum range of operation) and shocked with warm wair to heat up. Expect negative tending trift with 0°C tests.
UCT_AST9AH_B10 | -10 | 2024-05-15 09:30 | -9.6 |  | Between B01 and B06 | -9.8 | -8.4 | 9:54 | UCT_AST_B10_-10
UCT_AST9AH_A01 | 0 | 2024-03-08 08:00 | 0.5 | 8:16 | Between A01 and A06 | 0.5 | 3 | 8:16 | UCT_AST_9AH_A01_0 | Adjusted EIS program limits from 8kHz back to 5kHz. Arrived to discover all test units offline. Delayed while restoring power. Measurement nominal. | ETC sufered 2 power failures overnight but eithout significant deviation and normal operation for 10h before start of EIS measurements.
UCT_AST9AH_A02 | 0 | 2024-03-08 08:10 |  |  | Between A01 and A06 | 0.5 | 3 | 8:26 | UCT_AST_9AH_A02_0
UCT_AST9AH_A03 | 0 | 2024-03-08 08:20 |  |  | Between A01 and A06 | 0.5 | 3 | 8:36 | UCT_AST_9AH_A03_0
UCT_AST9AH_A04 | 0 | 2024-03-08 08:30 |  |  | Between A01 and A06 | 0.5 | 3 | 8:46 | UCT_AST_9AH_A04_0
UCT_AST9AH_A05 | 0 | 2024-03-08 08:40 |  |  | Between A01 and A06 | 0.5 | 3 | 8:56 | UCT_AST_9AH_A05_0
UCT_AST9AH_A06 | 0 | 2024-03-08 08:50 | 3 | 9:06 | Between A01 and A06 | 0.5 | 3 | 9:06 | UCT_AST_9AH_A06_0 |  | Ethyl Glycol measurement taken as A06 Removed
UCT_AST9AH_A07 | 0 | 2024-03-08 09:00 |  |  | Between A01 and A10 | 0.5 | 5.7 | 9:17 | UCT_AST_9AH_A07_0
UCT_AST9AH_A08 | 0 | 2024-03-08 09:10 |  |  | Between A01 and A10 | 0.5 | 5.7 | 9:27 | UCT_AST_9AH_A08_0
UCT_AST9AH_A09 | 0 | 2024-03-08 09:20 |  |  | Between A01 and A10 | 0.5 | 5.7 | 9:37 | UCT_AST_9AH_A09_0
UCT_AST9AH_A10 | 0 | 2024-03-08 09:30 | 5.7 | 9:47 | Between A01 and A10 | 0.5 | 5.7 | 9:47 | UCT_AST_9AH_A10_0
UCT_AST9AH_B01 | 0 |  | 0 | 14:08 | Between B01 and B06 | 0 | 3.1 | 14:08 | UCT_AST_B01_+00 | Nominal | Minimum chest freezer temperature is -5°C and prine to low swings. Elected to threefore use ETC as storage unit through EIS measurements, accepting upwards drift.
UCT_AST9AH_B02 | 0 |  |  |  | Between B01 and B06 | 0 | 3.1 | 14:18 | UCT_AST_B02_+00 | Sharp resonance in resistive axis in HF range.
UCT_AST9AH_B03 | 0 |  |  |  | Between B01 and B06 | 0 | 3.1 | 14:28 | UCT_AST_B03_+00 | Nominal
UCT_AST9AH_B04 | 0 |  |  |  | Between B01 and B06 | 0 | 3.1 | 14:38 | UCT_AST_B04_+00 | Nominal
UCT_AST9AH_B05 | 0 |  |  |  | Between B01 and B06 | 0 | 3.1 | 14:47 | UCT_AST_B05_+00 | Similar resonance as B02 but less extreme.
UCT_AST9AH_B06 | 0 |  | 3.1 | 14:57 | Between B01 and B06 | 0 | 3.1 | 14:58 | UCT_AST_B06_+00 | Nominal | Eth Glycol taken as B06 removed.
UCT_AST9AH_B07 | 0 |  |  |  | Between B01 and B10 | 0 | 6 | 15:07 | UCT_AST_B07_+00 | Minor resonance point near 3kHz
UCT_AST9AH_B08 | 0 |  |  |  | Between B01 and B10 | 0 | 6 | 15:17 | UCT_AST_B08_+00 | Nominal
UCT_AST9AH_B09 | 0 |  |  |  | Between B01 and B10 | 0 | 6 | 15:27 | UCT_AST_B09_+00 | Point of inflection in LF lobe.
UCT_AST9AH_B10 | 0 |  | 6 | 15:46 | Between B01 and B10 | 0 | 6 | 15:37 | UCT_AST_B10_+00 | Point of inflection in LF lobe. | All batteries in water bath. Began B01 at RT immediately after B10 at 0°C
UCT_AST9AH_A01 | 25 | 2024-03-08 13:00 |  |  | Room Ambiant Conditions | 25 | 27.2 | 13:13 | UCT_AST_9AH_A01_+RT2 | Minor water intrusions in top cover from water bath. Measurement nominal. Linear trend (Ferg suggests diffusion limit) now gone. | Placed in 25°C water bath to bring to room temperature. Then placed on shelf before EIS conducted. Assumed saturated, and room air temperature measurement assumed upper bound.
UCT_AST9AH_A02 | 26 | 2024-03-08 13:10 |  |  | Room Ambiant Conditions | 25 | 27.2 | 13:26 | UCT_AST_9AH_A02_+RT2 | Nominal
UCT_AST9AH_A03 | 25 | 2024-03-08 13:20 |  |  | Room Ambiant Conditions | 25 | 27.2 | 13:38 | UCT_AST_9AH_A03_+RT2 | Nominal
UCT_AST9AH_A04 | 25 | 2024-03-08 13:30 |  |  | Room Ambiant Conditions | 25 | 27.2 | 13:48 | UCT_AST_9AH_A04_+RT2 | Nominal
UCT_AST9AH_A05 | 25 | 2024-03-08 13:40 |  |  | Room Ambiant Conditions | 25 | 27.2 | 13:58 | UCT_AST_9AH_A05_+RT2 | Nominal
UCT_AST9AH_A06 | 25 | 2024-03-08 13:50 |  |  | Room Ambiant Conditions | 25 | 27.2 | 14:08 | UCT_AST_9AH_A06_+RT2 | Nominal | Placed A06-A10 into ETC at 25°C in order to dry out. Still wet when taken out however.
UCT_AST9AH_A07 | 25 | 2024-03-08 14:00 |  |  | Room Ambiant Conditions | 25 | 27.2 | 14:18 | UCT_AST_9AH_A07_+RT2 | Nominal
UCT_AST9AH_A08 | 25 | 2024-03-08 14:10 |  |  | Room Ambiant Conditions | 25 | 27.2 | 14:28 | UCT_AST_9AH_A08_+RT2 | Nominal
UCT_AST9AH_A09 | 25 | 2024-03-08 14:20 |  |  | Room Ambiant Conditions | 25 | 27.2 | 14:37 | UCT_AST_9AH_A09_+RT2 | Nominal
UCT_AST9AH_A10 | 25 | 2024-03-08 14:30 |  |  | Room Ambiant Conditions | 25 | 27.2 | 14:47 | UCT_AST_9AH_A10_+RT2 | Nominal. Battery appears recovered despite freezing prior,
UCT_AST9AH_B01 | 25 | 2024-05-15 15:45 |  |  | Room Ambiant Conditions | 27.8 | 27.8 | 15:47 | UCT_AST_B01_+RT2 | Nominal | Batteries dried and staged as prior completes EIS. After EIS B01-B08 placed back at -20°C for OCV test next morning on 16May. B01-B03 then placed in water bath to heat up (all else brought up in air).
UCT_AST9AH_B02 | 25 | 2024-05-15 15:55 |  |  | Room Ambiant Conditions | 27.8 | 27.8 | 16:38 | UCT_AST_B02_+RT2 | Accedentially ran B02X instead of B02. Corrected error in EIS comment. B02 ran after B05.
UCT_AST9AH_B03 | 25 | 2024-05-15 16:05 |  |  | Room Ambiant Conditions | 27.8 | 27.8 | 16:07 | UCT_AST_B03_+RT2 | Nominal
UCT_AST9AH_B04 | 25 | 2024-05-15 16:15 |  |  | Room Ambiant Conditions | 27.8 | 27.8 | 16:17 | UCT_AST_B04_+RT2 | Nominal
UCT_AST9AH_B05 | 25 | 2024-05-15 16:25 |  |  | Room Ambiant Conditions | 27.8 | 27.8 | 16:38 | UCT_AST_B05_+RT2 | Nominal
UCT_AST9AH_B06 | 25 | 2024-05-15 16:35 |  |  | Room Ambiant Conditions | 27.8 | 27.8 | 16:47 | UCT_AST_B06_+RT2 | Nominal | Placed momentrily in water bath but removed when audible noise heard.
UCT_AST9AH_B07 | 25 | 2024-05-15 16:45 |  |  | Room Ambiant Conditions | 27.8 | 27.8 | 16:57 | UCT_AST_B07_+RT2 | 1 errenous data point in LF range.
UCT_AST9AH_B08 | 25 | 2024-05-15 16:55 |  |  | Room Ambiant Conditions | 27.8 | 27.8 |  | UCT_AST_B08_+RT2 | HF range extremely polytomic. | Batteries B01-B08 placed back in ETC @ -20°C overnight for morning OCV test. B09, B10 left out for morning EIS.
UCT_AST9AH_B09 | 25 | 2024-05-15 17:05 |  |  | Room Ambiant Conditions | 27.8 | 27.8 |  | UCT_AST_B09_+RT2 | Failed EIS measurement repeatedly. Errenous points follow an (illogical) pattern, indicating issue in DSP. Eventually moved on to B10. B10 produced a successful result (after a 2nd attempt). B09 then attempted again and finally produced a nominal result.
UCT_AST9AH_B10 | 25 | 2024-05-15 17:15 |  |  | Room Ambiant Conditions | 27.8 | 27.8 |  | UCT_AST_B10_+RT2 | First measurement failed. Repeated EIS test and found nominal result. Slight cresting (nominal feature).
Adjustment: Changed from thermometer battery to 80v% Ethylene Glycol.
Adjustment: Performed initial PoC EIS Spectra at -20°C with reduced cooling period.
Delay: Only cooled to -40°C on end of Tuesday 5 March. Planned to then continue to time schedule as originally planned.
Issue: Advised that Ethylene Glycol will have a faster temperature response than batteries. Can be used to confidently state maximum drifts, assuming an initial thermally saturated condition.
ETC set to 0% relativehumidity throughout testing.
Added fan to cool room from heat emmitted from ETC while taking EIS measurements.
Modified methodology to place ETC setpoint -5°C below target to allow spanning of temperarture uncertainty across target.
Increased EIS range from 3kHz - 5kHz - 8kHz as real crossing lost in deeply discharged batteries.
4. Teardowns
Schedule
| Time per Teardown | 0:30
Only Record Density for 1 Cell per Battery
Hydrometer
Battery | Due | Call | Voltage | Density | Temperature | Notes / Observations
UCT_AST9AH_A01 | 2024-03-06 09:30 | I | 0 | 0 | 0
II
III
IV
V
VI
UCT_AST9AH_A02 | 2024-03-06 10:00 | I
II
III
IV
V
VI
UCT_AST9AH_A03 | 2024-03-06 10:30 | I
II
III
IV
V
VI
UCT_AST9AH_A04 | 2024-03-06 11:00 | I
II
III
IV
V
VI
UCT_AST9AH_A05 | 2024-03-06 11:30 | I
II
III
IV
V
VI
UCT_AST9AH_A06 | 2024-03-06 12:00 | I
II
III
IV
V
VI
UCT_AST9AH_A07 | 2024-03-06 12:30 | I
II
III
IV
V
VI
UCT_AST9AH_A08 | 2024-03-06 13:00 | I
II
III
IV
V
VI
UCT_AST9AH_A09 | 2024-03-06 13:30 | I
II
III
IV
V
VI
UCT_AST9AH_A10 | 2024-03-06 14:00 | I
II
III
IV
V
VI
UCT_AST9AH_B01 |  | I
II
III
IV
V
VI
UCT_AST9AH_B02 |  | I
II
III
IV
V
VI
UCT_AST9AH_B03 |  | I
II
III
IV
V
VI
UCT_AST9AH_B04 |  | I
II
III
IV
V
VI
UCT_AST9AH_B05 |  | I
II
III
IV
V
VI
UCT_AST9AH_B06 |  | I
II
III
IV
V
VI
UCT_AST9AH_B07 |  | I
II
III
IV
V
VI
UCT_AST9AH_B08 |  | I
II
III
IV
V
VI
UCT_AST9AH_B09 |  | I
II
III
IV
V
VI
UCT_AST9AH_B10 |  | I
II
III
IV
V
VI
Formulas:
Task = [Task, row above]+C101

## Sheet1
Battery | Estimated SoC | OCV
UCT_AST9AH_B01 | 100% | 12.78
UCT_AST9AH_B02 | 93.33% | 12.88
UCT_AST9AH_B03 | 86.67% | 12.61
UCT_AST9AH_B04 | 80% | 12.46
UCT_AST9AH_B05 | 73.33% | 12.32
UCT_AST9AH_B06 | 66.67% | 12.21 | 11v5 appeared initially
UCT_AST9AH_B07 | 60% | 12.06
UCT_AST9AH_B08 | 53.33% | 12.04
UCT_AST9AH_B09 | 46.67% | 12
UCT_AST9AH_B10 | 40% | 12.01
B01-B08 Eth Glycol | -19.4 | 2024-05-16 08:15
B09-B10 Eth Glycol | -19.3 | 2024-05-16 16:15